## Slide 1: February 2017 Øystein Prytz
MENA3100
Spectroscopy in the TEM and SEM

## Slide 2: Curriculum
Chapter 6 of the textbook
“Spectroscopy with incident x-rays and with incident electrons” by Johan Taftø
This lecture
Main focus is EDS in SEM and TEM, we also cover basic aspects of EELS

## Slide 3: Simple overview of the TEM

## Slide 4: Electron beam – sample interaction
W&C

## Slide 5: Energy transfer processes
The main energy transfer processes are:
Single electron excitations
Bremsstrahlung
Collective excitations (plasmons)
The first two processes are observed both in energy dispersive X-ray spectroscopy (EDS) and in electron energy loss spectrscopy (EELS).
The last process is only observed in EELS.

## Slide 6: Single electron excitations
1: Incident electron transfers energy to a «core» electron of the atom, which exits the sample.
2: The atom is left in an excited state with a core hole.
3a: The atom can relax to the ground state either by emitting a photon (characteristic X-ray)…
3b:…or by emitting an Auger-electron

## Slide 7: Nomenclature
Atomic shell | Main quantum number n | Orbital quantum number l | Total spin quantum number j | Spectroscopic notation
1s | 1 | 0 | +1/2, -1/2 | K
2s | 2 | 0 | +1/2, -1/2 | L1
2p1/2 | 2 | 1 | 1/2 | L2
2p3/2 | 2 | 1 | 3/2 | L3
3s | 3 | 0 | +1/2, -1/2 | M1
3p1/2 | 3 | 1 | 1/2 | M2
3p3/2 | 3 | 1 | 3/2 | M3
3d3/2 | 3 | 2 | 3/2 | M4
3d5/2 | 3 | 2 | 5/2 | M5

## Slide 8
Nomenclature
W&C

## Slide 9: More on nomenclature
W&C
Siegbahn notation is most commonly used to name the transitions generating X-rays.
However, this can get quite complicated as seen in the figure on the left.
The International Union of Pure and Applied Chemistry (IUPAC) recommends an alternative system that is simpler, but unfortunately not widely in use.

## Slide 10
R. Jenkins et al. X-Ray Spectrometry 20(3), 149-155 (1991).

## Slide 11: Single electron excitations
1: Incident electron transfers energy to a «core» electron of the atom, which exits the sample.
2: The atom is left in an excited state with a core hole.
3a: The atom can relax to the ground state either by emitting a photon (characteristic X-ray)…
3b:…or by emitting an Auger-electron

## Slide 12: Copper K lines

## Slide 13: Copper K lines
Eb(K) = -8979 eV
Eb(L3) = -933 eV Eb(L2) = -952 eV
Eb(M2,3) = -76 eV
L3 -> K transition (Kα1)
E=hn= Eb(L3)- Eb(K) = 8.046 keV
L2 -> K transition (Kα2)
E=hn= Eb(L2)- Eb(K) = 8.027 keV
M2,3 -> K transition (Kb)
E=hn= Eb(M2,3)- Eb(K) = 8.903 keV

## Slide 14: Threshold/critical energy
In order to generate X-rays, the electron beam must have an energy E0 larger than the critical energy Ec of the process.
Usually not a problem in TEM
E0> 100 keV; Ec< 20 keV
BUT: with Cs correctors, low voltage operation has become more common. 60 keV, 40 keV, even down to 30 keV.
For heavy elements this may limit which characteristic X-rays are generated

## Slide 15: Single electron excitations
1: Incident electron transfers energy to a «core» electron of the atom, which exits the sample.
2: The atom is left in an excited state with a core hole.
3a: The atom can relax to the ground state either by emitting a photon (characteristic X-ray)…
3b:…or by emitting an Auger-electron

## Slide 16: The fluorescence yield
The probability for generating a characteristic X-ray is given by the fluorescence yield w
The probability of generating an Auger electron is the 1- w.
F&H

## Slide 17: Energy transfer processes
The main energy transfer processes are:
Single electron excitations
Bremsstrahlung
Collective excitations (plasmons)
The first two processes are observed both in energy dispersive X-ray spectroscopy (EDS) and in electron energy loss spectrscopy (EELS).
The last process is only observed in EELS.

## Slide 18: Bremsstrahlung
The change in trajectory is caused by an acceleration of the electron. The spectrum of the photons generated is given by the Fourier transform of a(t) for the electron.
F&H

## Slide 19: Bremsstrahlung
Single bremsstrahlung event
Several events
Many events
Kramer’s cross section for production of bremsstrahlung
F&H
W&C

## Slide 20: Instrumentation
How ca we detect detect X-rays and measure their energy?

## Slide 21: The detector and electron/hole-pair generation
Characteristic X-rays are generated in the specimen and enter the detector
There, they generate a number of electron/hole-pairs depending on their energy
The electron/hole-pairs are separated by an applied bias, and the current measured
W&C

## Slide 22: The detector and electron/hole-pair generation
Historically, the most common detectors have been the Si(Li) detectors
The so-called Silicon-drift detectors (SDD) have been used for SEM for quite a while, and are becoming more common also for TEM
The detectors are usually cooled to avoid
  diffusion of dopants in the strong applied bias
  thermal noise
Si(Li) detectors are usually cooled to LN2 temperatures (-196 C)
SDD detectors can make do with only Peltier cooling (~-30 C)
W&C

## Slide 23: WDS vs EDS
Uses a diffraction grating (crystal) to select X-rays with particular energy (wavelength)
Braggs law
Wavelength dispersive, in stead of energy dispersive
Excellent energy resolution
Low background, no detector artefacts
Good for light elements
Serial detection
Movable parts
Low effective detection angle

## Slide 24: Quantification from EDS spectra
How to get from a spectrum to composition
Assumptions usually made in EDS in TEM
Cliff-Lorimer k-factor method
Limits of the CL-method and the assumptions made

## Slide 25: How to get from a spectrum to composition

## Slide 26: X-ray generation in the SEM
F&H

## Slide 27: ZAF and z [phi-rho-z]
Z – how the electrons penetrate the sample (stopping power and backscatter)
A – Absorption of X-rays generated
F – Fluorescence of secondary photons
z is the function describing the «mass-depth» dependence of X-ray production

## Slide 28: The thin film approximation
F&H
In the SEM
In the TEM

## Slide 29: The Cliff-Lorimer equations
F&H
Binary system
Ternary system

## Slide 30: k-factors are not constants
W&C

## Slide 31
Sheridan (1989)

## Slide 32: Limits to the thin film approximation
F&H

## Slide 33: Limits to the thin film approximation
F&H

## Slide 34: Single electron excitations
1: Incident electron transfers energy to a «core» electron of the atom, which exits the sample.
2: The atom is left in an excited state with a core hole.
3a: The atom can relax to the ground state either by emitting a photon (characteristic X-ray)…
3b:…or by emitting an Auger-electron

## Slide 35: The post column Electron Energy Loss spectrometer/energy filter (EELS)
D. B. Williams and C. B. Carter (2009)

## Slide 37: The magnetic prism
The force experienced by the electrons is a function of velocity (energy)
Different velocities give different curvature of the path
Non-relativistic case!
Increasing energy loss

## Slide 38: Sketch of an EELS spectrum
Zero energy loss peak (ZLP, 0 eV)
Low loss region. Valence band excitations, dominated by plasmons (collective excitations), dielectric properties ( 0-50 eV)
Core loss region, single electron transitions, bonding and density of states, progressively weaker signals (50 - 2000 eV)
R. Brydson (2001)

## Slide 39: Inner shell edges (core loss edges)
unoccupied states
occupied states
2p states
3s states
3p states
3d states
4s states
3d states
Energy
eF
Incident TEM electron
Transmitted electron with energy loss

## Slide 40
Valence band
Conduction band
«Core» electrons
«Core» electrons
«Core» electrons
«Core» electrons
Energy
eF
Incident TEM electron
Transmitted electron with energy loss
Reminder:
1s: K shell
2s: L1 shell
2p: L2,3 shell
3s: M1 shell
3p: M2,3 shell
3d: M4,5 shell
…

## Slide 42: Uses of the core-loss features
Composition analysis and mapping (EFTEM)
Electronic structure and bonds (ELNES), comparison with DOS
Structure and atomic coordination (EXELFS)

## Slide 43: Interband transitions, band gaps
Valence band
Conduction band
«Core» electrons
«Core» electrons
«Core» electrons
«Core» electrons
Energy
eF
The smallest possible energy loss corresponds to the band gap1
1Except phonons, Cerenkov etc

## Slide 44: Uses of the low-loss features
Phase sensitive mapping
Thickness measurements and mapping
Investigations of valence electron density variations (change in plasmon energy)
Band gaps and interband transitions
Optical properties (dielectric function, surface plasmons)

## Slide 45: The low loss region
The low loss region contains both collective excitations and single electron events
Plasmons
Valence band/interband excitations (VEELS)
Also sometimes «core» loss events

## Slide 46: Collective oscillations of the valence electrons - plasmons
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
e-
Characteristic frequency wp and energy Ep
Not atom characteristic, but phase characteristic

## Slide 47: Comparison of geometry of EDS and EELS
The EDS detector has a fixed solid angle, and detect only a small fraction of the generated X-rays
The collection angle  of the EELS detector can be varied, and often captures the majority of electrons that have undergone energy loss

## Slide 48: Energy resolution
FWHM = 1.6 eV
Source | Minimum energy spread (eV FWHM)
W (thermionic) | ~3 eV
LaB6 (thermionic) | ~1.5 eV
Thermally assisted field emission (Schottky FEG) | ~0.7 eV
«Cold» FEG | ~0.3 eV
Monochromated FEG* | 0.1-0.3 eV
* No fundamental limit ? | 
The energy resolution is usually measured as the Full Width at Half Maximum (FWHM) of the Zero Loss Peak (ZLP)
The main determining factor is usually the energy spread of the electron source
However, other factors also have a strong influence
  Voltage stability
  Spectrometer stability
  Magnetic fields

## Slide 49
FWHM = 1.6 eV
FWHM = 0.15 eV

## Slide 51: Wien filter monochromator (FEI)
W. Grogger et al. Top. Cat. (2008)
Electric and magnetic fields tuned so that the force on the electrons balance out for one specific velocity (energy) v0.
Electrons with this velocity continue straight ahead, other are spatially distributed
Electron beam after monochromator is spatially dispersed according to energy

## Slide 52: Effect of the improved energy resolution

## Slide 53: Summary
The main energy transfer processes are:
  Single electron excitations
  Bremsstrahlung
  Collective excitations (plasmons)
Differentiate between primary and secondary processes
Nomenclature
What determines the characteristic X-ray energies
How can X-ray energies be measured?
  EDS vs WDS
The thin film approximation in TEM
  What approximationis made?
  What are the limits to this approximation and what happens when it breaks down?
  Difference between the Cliff-Lorimer method and the ZAF method
Electron Energy Loss Spectroscopy (EELS)
  The EELS spectrometer
  Typical features in the EELS spectrum and their uses
  Monochromation