## Slide 1
MEK400 – Experimental methods in fluid mechanics
Introduction to Particle Image Velocimetry (PIV)
26.2.2013
J. Kristian Sveen (IFE/FACE/UiO)

## Slide 2: This presentation looks at how to use pattern matching to measure velocities
Pattern matching in PIV
Challenges – solutions
Laboratory application
  Seeding,
  illumination,
  imaging

## Slide 3: The human brain is great at matching patterns
Computers perhaps a little less great

## Slide 4: Pattern matching in everyday applications
Locating a face in an image
Identifying a number plate on a car
Finding motion of random patterns

## Slide 5: Pattern matching in PIV
Two consecutive images with known time spacing
Match pattern locally between corresponding grid cells
Divide into grid

## Slide 6: Pattern matching principles is the foundation for PIV
t1
t2

## Slide 7: The principle of Pattern Matching in PIV is to measure similarity of a local pattern in two subsequent images
Distance Metrics:
In which overlapping position are two images
  The most alike?
  The least different?
(any) introductory book on image processing will point to
CROSS CORRELATION:

## Slide 8: Cross correlation is a simple measure of similarity
For each sub-window pair
overlay sub-windows in all possible combinations
Matlab example (corrshifter.m)

## Slide 9: Cross correlation may easily be calculated using FFT’s
Sensitive to:
Amplitude change
Background gradients
Finite images (edge effects)
…
Correlation theorem (look it up)

## Slide 10: Sensitivity of cross correlation to image features
Amplitude – What happens if intensity in f is doubled from t1 to t2?
Background – What happens if background is non-zero and non-uniform?

## Slide 11: Removing effects of background
Subtract background from f and g before calculating correlation
Correlation signal including background
Correlation signal with background removed

## Slide 12: Normalization of correlation signal
Assuming means have been subtracted
Common simplification assumes evenly distributed pattern (standard deviation does not change locally):

## Slide 13: Correcting for loss of pattern
If pattern moves “many pixels” between frames information is lost
Only a part of the window (pattern) contributes to correlation signal
Same applies for large velocity differences across windows
Leads to a bias towards smaller values (see Westerweel, 1993)
Use window shifting to improve correlation

## Slide 14: Sub-pixel displacement estimation
By interpolating the peak in the correlation plane, sub-pixel accuracy may be achieved.

## Slide 15: Peak interpolation
3 common interpolation schemes
  Center of mass
  Parabolic fit
  Gaussian fit
R0
R-1
R+1

## Slide 16: When the peak becomes narrow, sub-pixel resolution may be lost
May lead to “peak-locking”
-only the central lobe contributes

## Slide 17: Also the interpolation scheme may contribute to peak locking
The traditional solution is to use sub-pixel window shifting
Requires substantial image interpolation and iteration
error

## Slide 18: What happens in regions with background gradients?
Standard FFT based correlation
Background gradients have huge influence on result

## Slide 19: Our image example
Our standard FFT based correlation
The correct peak
…a few other correlation functions

## Slide 20: Vector validation
Our vector field…
Clearly some vectors are wrong?
How do we determine this?

## Slide 21: Vector validation – global view
Identify vectors that are significantly different from average
plot u vs v
Drawback: if mean is used, faulty vectors contribute to the mean

## Slide 22: Vector validation – local view
Use smaller regions for comparison
If vector is significantly different from 8 or 24 neighbors – it may be discarded
Use mean or median:
Median safer – less likely to be biased by the faulty vector(s)

## Slide 23: Vector validation – signal to noise ratio
Compare peak height to second highest peak in correlation plane
Quality of signal compared to level of noise
Often also referred to as a detectability measure

## Slide 24: “Alternative” correlation functions
Often referred to as “Distance metrics”
Minimum quadratic difference (Gui&Merzkirch,2000):
Recognise this?

## Slide 25: “Alternative” correlation functions
Normalised correlation is often a better choice over standard FFT based correlation since it handles pattern variation better

## Slide 26: “Alternative” correlation functions
Looking back at the FFT based correlation:
If amplitude variations hamper the precision – is it possible to reduce the effect by, say, using Phase correlations?
Removing the amplitude works, but we loose precision

## Slide 27: Phase correlations in PIV
Phase correlations have been applied in PIV by several authors due to robustness to noise
Use as a first iteration step
Phase corr
mqd

## Slide 28: A short summary

## Slide 29: PIV in the laboratory

## Slide 30: The practical aspects of PIV
So far: software principles
From www.dantecdynamics.com
Next: what we do in the laboratory

## Slide 31: Seeding of flow
For pattern matching to work, we need
A pattern
Images of the pattern
Ludwig Prandtl used particles in visualization experiments in the 1920’s and 1930’s
- Small aluminum particles
See www.dlr.de

## Slide 32: Types of seeding material
 | PSP Polyamide seeding particles | HGS Hollow glass spheres | S-HGS Silver-coated hollow glass spheres | FPP Fluorescent polymer particles
Mean particle size (µm) | 5, 20, 50 | 10 | 10 | 10, 30
Size distribution | 1 - 10 µm 5 - 35 µm 30 - 70 µm | 2 - 20 µm | 2 - 20 µm | 1 - 20 µm 20 - 50 µm
Particle shape | non-spherical but round | spherical | spherical | spherical
Density (g/cm3) | 1.03 | 1.1 | 1.4 | 1.19
Melting point (°C) | 175 | 740 | 740 | 125
Refractive index | 1.5 | 1.52 | — | 1.479
Material | Polyamide 12 | Borosilicate glass | Borosilicate glass | Poly (Methyl methacrylate)(Labeled with Rhodium B)
Requirement: passive tracers that follow the flow
Dust, smoke, aerosols, dirt, pollen, chemicals
- Anything that forms a pattern

## Slide 33: Size of seeding particles
From the software side: particles need to cover more than ~2.35 pixels (diameter) to limit peak-locking errors
From the experimental side: how closely does the particle velocity V follow the fluid velocity v?
Compare slip velocity |v-V| to stokes drag on a sphere

## Slide 34: Particle sizes
T=5-10s, n=10-6, R=0.5mm
0.5-1% error

## Slide 35: Imaging
We need to accurately acquire two consecutive images with a known time spacing
With a 10cmx10cm imaging area (Field of View), imaged by a camera with 1000x1000 pixels, implies 100 pixels per centimetre.
A flow of just 10cm/second = 1000 pixels per second
To recover this in a 32x32 interrogation window, the pattern should ideally move less than 16 pixels (why?)
 16p / 1000p/s = 16mseconds between frames  62.5 frames per second (if regular camera)

## Slide 36: Imaging – types of cameras
Special purpose PIV cameras often used
Trigger by dual-cavity laser at end of frame 1 and start of frame 2
Very low interframe times possible (nanoseconds)
Alternative: high speed cameras (~7000 fps @ megapix resolution)

## Slide 37: Calibration from pixels to centimeters
We need to convert from pixels to centimeters
Solution: image a grid with known spacing
Simple convertion
XX pixels = YY centimeters

## Slide 38: Writing your own PIV code
Simple PIV